## Slide 1: Progress of Injector Linac
Kazuro Furukawa
for Injector Linac

## Slide 2
概要
Alignment
電子銃
RF 電子銃
熱電子銃
陽電子発生装置
Schedule

## Slide 3: Linac Upgrade Overview

## Slide 4: Mission of electron/positron Injector in SuperKEKB
40-times higher Luminosity
  Twice larger storage beam  Higher beam current at Linac
  20-times higher collision rate with nano-beam scheme
     Low-emittance even at first turn  Low-emittance beam from Linac
     Shorter storage lifetime  Higher Linac beam current
Linac challenges
  Low emittance e-
    with high-charge RF-gun
  Low emittance e+
    with damping ring
  Higher e+ beam current
    with new capture section
  Emittance preservation
    with precise beam control
  4+1 ring simultaneous injection
Linac Upgrade Overview
Notes: Higher Injection Beam Current
  To Meet the larger stored beam current and shorter beam lifetime in the ring
  4~8-times larger bunch current for electron and positron
  Reconstruction of positron generator, etc
Lower-emittance Injection Beam
  To meet nano-beam scheme in the ring
  Positron with a damping ring
  Electron with a photo-cathode RF gun
  Emittance preservation by alignment and beam instrumentation
Quasi-simultaneous injections into 4 storage rings
  SuperKEKB e–/e+ rings, and light sources of PF and PF-AR
  Improvements to beam instrumentation, low-level RF, controls, timing, etc

## Slide 5: 電子ビームパラメタ
 | SuperKEKB | KEKB
エネルギー (GeV) | 7.0 | 8.0
HER蓄積電流値 (A) | 2.6 | 1.1
HERビーム寿命 (min.) | 6 | 200
最大ビーム繰り返し (Hz) | 50 | 50
最大バンチ数 (rfパルス当たり) | 2 | 2
エミッタンス (mmmrad) | 50/20 (Hor./Ver.) | 100
バンチ電荷量 (nC) | 5 | 1
エネルギー広がり (%) | 0.1 | 0.05
バンチ長 sz (mm) | 1.3 | 1.3
ダンピングリング | n/a | n/a
同時トップアップ入射 | 4 rings (SuperKEKB e-/e+, PF, PF-AR) | 3 rings (KEKB e-/e+, PF)
Linac Upgrade Overview

## Slide 6: 陽電子ビームパラメタ
 | SuperKEKB | KEKB
エネルギー (GeV) | 4 | 3.5
LER蓄積電流値 (A) | 3.6 | 1.6
LERビーム寿命 (min.) | 6 | 133
最大ビーム繰り返し (Hz) | 50 | 50
最大バンチ数 (rfパルス当たり) | 2 | 2
エミッタンス (mmmrad) | 100/20 (Hor./Ver.) | 2100
バンチ電荷量 (nC) | 4 | 1
エネルギー広がり (%) | 0.1 | 0.125
バンチ長 sz (mm) | 0.7 | 2.6
ダンピングリング | 〇 | n/a
同時トップアップ入射 | 4 rings (SuperKEKB e-/e+, PF, PF-AR) | 3 rings (KEKB e-/e+, PF)
Linac Upgrade Overview

## Slide 7: Linac Schedule Overview
RF-Gun e- beam
commissioning
at A,B-sector
e- commiss.
at A,B,J,C,1
e+ commiss.
at 1,2 sector (FC, DCS, Qe- 50%)
e- commiss.
at 1,2,3,4,5 sector
non damped e+ commiss.
at 1,2, 3,4,5 sectors
e- commiss. at A→5 sectors
damped e+ commiss.
at 1→5 Qe+ = 1~4nC
e- commiss.
at A→5 Qe- = 1~5nC
: Electron
: Positron
: Low current electron
2 nC
Phase2
1 nC
Phase1
Location
↓
Time →
Phase1: high emittance beam for vacuum scrub
Phase2,3: low emittance beam for collision
Schedule
Without Top-up
4 nC
Phase3
Improved RF gun
Direct PF-AR BT

## Slide 8: Operation Status
FY2014 の運転
  運転時間: 3900 hours (FY2013 -27%)
    PF/PF-AR Injection (and Commissioning for SuperKEKB)
    Apr.11 – Apr.25
    May.7 – Jul.1
    Sep.24 – Dec.26
    5 days in Jan, Feb
  故障 (FY2013 0.43%)
    偏向電磁石コイル発熱 (May)
    マイクロ波パルス電源デスパイカー発熱 (May)
    フラックスコンセントレータ電源側ケーブル発熱 (Dec)
    マイクロ波パルス電源トランス端末発熱 (Dec)
Operation

## Slide 9: Facility for Electric Power and Cooling Water
4.0 GeV e+
4nC x 2
3.5 GeV
10nC x 2 (prim. e-)
5nC x 2 (inj. e-)
1.1 GeV
Damping Ring
circ. 136 m
ECS
Energy-spread
Compression System
Bunch
Compressor
7.0 GeV e-
5nC x 2
low emittance
RF gun
SuperKEKB injector
e+ target &
LAS capture section
50 Hz (e+ or e-)
pulse-by-pulse
mode switching
S-band
linac
2.5 GeV e-
0.1nC x 1
PF
HER
LER
AR
6.5 GeV e-
3T
RF gun
[1] Sector 1,2
Electricity and cooling water addition for new positron generator
+920 kVA, +900 L/min
[2] Sector 2,3
Electricity addition for new damping ring
+310 kVA
[3] Sector 5
Electricity and cooling water addition for energy compression, beam diag., injection beamlines
+300 kVA, +460L/min
Linac needs electricity and cooling water extensions, especially for positron generator upgrade
Separate building construction in FY2013 not to impact PF/PF-AR
Facility extensions performed during summer 2014
Linac Upgrade Overview

## Slide 10: Linac Upgrade Progress towards SuperKEKB (1)
High-charge low-emittance RF gun development
  QTWSC cavity and Ir5Ce photo cathode works well
Positron generation confirmation for the first time
  Good agreement with the simulation results
Precise alignment for emittance preservation
  Recovering after earthquake
  Reaching specification of 0.3mm
Utility upgrade during summer 2014
  for electricity (+1.5MW) and cooling water (+1400L/min)
Signal from primary electron
Signal from positron with opposite polarity
Ir5Ce Cathode
Quasi traveling wave side couple cavity
Positron generator
Linac Upgrade Overview

## Slide 11: Linac Upgrade Progress towards SuperKEKB (2)
High power modulator upgrades
Low-level RF controls/monitor
  Pulse-to-pulse modulation (PPM) between 4+1 rings
  More spaces for increased number of devices
Beam instrumentation
  Large/small aperture beam position monitors (BPM)
  Precise/fast and synchronized BPM readout system
  Wire scanners and beam loss monitors
  Streak cameras
  (Deflectors, etc)
Event timing control system
  Combination of MRF and SINAP modules
  Essential for PPM operation
  Precise timing & synchronized controls
  Bucket selection at DR and MR
SINAP event modules
Linac Upgrade Overview
Beam wire scanner

## Slide 12: Alignment
Alignment

## Slide 13: Alignment
High-precision alignment was not necessary in PF and KEKB injections, and it was much damaged by earthquake in 2011.
Instead of flexible-structure girder before earthquake, rigid-structure was adopted with jack-volts and fixed supports.
Reflector pedestals are developed and mounted onto quad magnets and accelerating cavities for laser-tracker measurement.
Iterative measurement and adjustment with 500-m straight laser and position sensors should enable 0.3-mm global alignment.
Laser tracker should enable 0.1-mm measurement within 10-m girder unit.
Displacement gauges, hydrostatic leveling, inclinometer are also employed.
Remote measurement system and girder mover system will be necessary for longer term, and are under development.
Alignment

## Slide 14: Emittance Preservation and Alignment
If Device is off center of the beam
  Focusing magnet (quad) kicks the beam bunch
  Accelerating structure (cavity) excites wakefield, to bend the tail
Distorted bunch in banana shape
  Emittance dilution or blow-up, even 100 times larger
    Depending on the beam optics and the beam charge
Alignment and orbit correction is crucial to preserve the emittance
Focusing Magnet
Beam
Accelerating Structure
Transverse beam distribution in time direction
Sugimoto et al.
Alignment

## Slide 15: Emittance Preservation
Offset injection may solve the issue
Orbit have to be maintained precisely
Mis-alignment should be <0.1mm locally, <0.3mm globally
100 samples
Mis-alignment leads to Emittance blow-up
Sugimoto et al.
Orbit manipulation compensates it
Alignment

## Slide 16: Reflector setting for hardware positioning
Mounted on coupler OD
Arm from Laser PD system to reflector
Reflectors for old Q-magnet
Higo et al.
Alignment

## Slide 17: Alignment progress in 2014
For the first time at 3-5 sectors
Horizontal axis: sensor number from sector C, 1-4, to sector 5 (~80 m/sector)
Vertical axis: voltage (~displacement at 0.25~0.5mm/mV)
Some girders were not yet upgraded
Moved up to 3 mm not to break vaccuum
Alignment
Displacement / mV
Sector C
Sector 5

## Slide 18: Alignment progress in 2014
For the first time after earthquake at downstream sectors
Several measurements during summer
Measurement reproducibility was confirmed up to ~0.2 mm
While there existed several conflicting measurements, consistent scheme has been established
Movement of tunnel by several 10’s of micrometer was observed (→ mover)
Further work necessary in 2015, for alignment and girder replacement
Higo et al.
~5mm
~10mm
Alignment

## Slide 19: Recent 500-m alignment over C-5 after completion of initial alignment in late Jan. 2015
Higo et al.
Alignment

## Slide 20: Hardware alignment on girders in sectors 3~5
Horizontal
s=34mm
Vertical
s=47mm
Horizontal
Vertical
Higo et al.
Alignment

## Slide 21: Floor horizontal movement in a half year from summer to winter
C3
11
18
28
38
48
Higo et al.
Alignment

## Slide 22: Floor vertical movement in a half year from summer to winter
C3
11
18
28
38
48
Higo et al.
Alignment

## Slide 23: Junction relative movement at 28 19 Sep. 2014 – 31 Jan. 2015
Up
North (beam)
East
↑ Up, East, North(beam line)
↓ Down, West, South
23 Nov.
1 Feb.
Air cond. restart
Water flood due to typhoon
0.4
0.3
0.2
0.1
0.0
-0.1
-0.2
Linac stop
14 Sep.
Found daily movement, month-long drift, climate effect, etc. These are related to those observed in laser PD long-term result.
Higo et al.
Alignment

## Slide 24: Estimation of floor movement some typical observed values
 | Horizontal | Vertical
Daily | 0.1 | 0.1
Week | 0.1 | 0.1
Half a year | 0.5 | 2
Speed | 0.01mm/hour | 0.01mm/hour
Unit mm
We should study/develop the linac system
with these values in mind.
Precise beam orbit control is necessary to preserve emittance
Alignment

## Slide 25: RF gun

## Slide 26: RF-Gun development strategy for SuperKEKB
Cavity : Strong electric field focusing structure
  Disk And Washer (DAW) => 3-2, A-1(test)
  Quasi Traveling Wave Side Couple => A-1
  => Reduce beam divergence and projected emittance dilution
Cathode : Long term stable cathode
  Middle QE (QE=10-4～10-3 @266nm)
  Solid material (no thin film) => Metal composite cathode
  => Started from LaB6 (short life time) => Ir5Ce has very long life time and QE>10-4 @266nm
Laser : Stable laser with temporal manipulation
  LD pumped laser medium => Nd / Yb doped
  Temporal manipulation => Yb doped => Minimum energy spread
RF gun for low-emittance electron

## Slide 27: Photo cathode RF gun development
5.6 nC / bunch was confirmed
Next step: 50-Hz beam generation & Radiation control
Quasi traveling wave side couple cavity
Cascaded frequency doublers
Yb fiber and Yb:YAG think disk laser
Ir5Ce photo cathode
Ir5Ce Cathode
RF gun for low-emittance electron

## Slide 28: Schedule
These few months:
  Reconfigure thermal gun for positron generation
  Step by step RF-Gun RF ageing.
  New laser system in ground laser room.
    Increase pulse energy from fiber laser
    Simplify the laser system (new multi-pass amplifier)
This comming summer:
  Third RF-Gun (nomal laser injection / cavity modification / cathode change)
  Simple Nd amplifier for Phase-I & II stable injection
10 ps gaussian is enough for Phase-I (1nC) & II (2nC), postponed the pulse shaping until Phase-III
Following the RF-Gun reviewer’s comments
(Postponed the pulse shaping until Phase-III )
Yoshida et al.

## Slide 29
Thermionic DC gun
RF gun
SHB1
SHB2
prebuncher
buncher
Acc.tube
Acc.tube
Energy slope
30 MeV
20 psec
Chicane
Bunch compression
20 psec
triplet
5 nC
Low emittance
10 nC
for positron primary
Beam line will be upgrade on up and down.
Thermionic DC gun will be installed to upper beam line.
70 MeV
Yoshida et al.

## Slide 30: Reconfiguration of A-1 injector area
KL_A1_A
KL_A1_B
RF-Gun1
RF-Gun2
加速管

## Slide 31
Second RF gun on the 45 degree line
Studies on
Improved RF-Gun cavity
Normal laser injection
Cathode change including alkaline cathode
Normal laser injection
Angled laser injection

## Slide 32
Second Side coupled Quasi-travelling wave RF-Gun
Conditioning progress was too slow. Frequent break down is the issue to be cured.
Cathode rod contact?
Cathode material fixation?
Cathode material sputtering due to laser?
We have to analyze break-down issues.
Cavity conditioning, used dummy cathode rod without cathode material (all Cu).
Replace new cathode rod with material (new fixation is shrinkage fit).
For reduce multipactoring effect, another cathode cell design is required.
Second RF-Gun under brazing
shrinkage fit
Present cathode
New cathode

## Slide 33: Short term plan for laser development
Undeground Yb-Fiber + Yb:YAG (Existing)
  Downgrade to 25Hz 2-loop amplifier(Done p.8) => Fix configuration
  Increase monitor points / quadrant detector etc..
  Improve stability.
Underground Yb-Fiber + Nd:YAG (Yb-Fiber small upgrade)
  1064nm(Nd:YAG wavelength) is converted by existing Yb oscillator. (p.9)
  Stretcher for 10 ps is similar to existing one. (p.11)
  Existing fiber amplifier (Thorlabs) is best fit to amplify 1064nm. (p.9)
  Preliminary test using existing Nd:YAG DPSS Module (10Hz). (p.10)
Ground Yb-Fiber(Commercial) + Cryogenic Yb:YAG => Postpond the operation until Phase-III.
Ground Yb-Fiber(Commercial) + Nd:YAG(Commercial)
  MENLO Orange oscillator wavelength must be shifted. (p.9) or use Nd:YLF (1047nm). (p.9)
  Yb-Fiber commercial amplifier can be used for 1064nm. (p.10)
  Nd:YAG 50Hz DPSS commercial module (p.13)
  Vacuum duct / Room environment / Virtual cathode.

## Slide 34: Short term plan for laser development
Following recommendations at review meetings
Undeground Yb-Fiber + Yb:YAG (Existing)
  Downgrade to 25Hz 2-loop amplifier(Done) => Fix configuration
  Increase monitor points / quadrant detector etc..
  Improve stability.
Underground Yb-Fiber + Nd:YAG (Yb-Fiber small upgrade)
  1064nm(Nd:YAG wavelength) is converted by existing Yb oscillator.
  Stretcher for 10 ps is similar to existing one.
  Existing fiber amplifier (Thorlabs) is best fit to amplify 1064nm.
  Preliminary test using existing Nd:YAG DPSS Module (10Hz).
Ground Yb-Fiber(Commercial) + Cryogenic Yb:YAG => Postponed the operation until Phase-III.
Ground Yb-Fiber(Commercial) + Nd:YAG(Commercial)
  MENLO Orange oscillator wavelength must be shifted. or use Nd:YLF (1047nm).
  Yb-Fiber commercial amplifier can be used for 1064nm.
  Nd:YAG 50Hz DPSS commercial module
  Vacuum duct / Room environment / Virtual cathode.

## Slide 35
New high gain multi-pass amplifier(10-15 pass x 2 loop)
to simplify the laser
LD
Laser Diode
INPUT
OUTPUT
1pass
10-15pass
←

## Slide 36: Spectrum

## Slide 37: Simplify and stabilize our laser system without pulse shaping
Yb-doped fiber oscillator with Transmission grating
PCF Yb-doped fiber amplifier
Pulse picker
Commertial Nd:YAG rod
multipass amplifier (10 Hz => 50 Hz)
Yb:YAG thin-disk multipass amplifier
(25 Hz)
114.2 MHz
@ 1030 nm
@ 1064 nm
LIEKKI Yb-doped fiber amplifier
Pulse picker
Grating pair
Fiber Amplifier @1030 nm
Fiber Amplifier @1064 nm
Spatially dispersed
Yb fiber Oscillator
Menlo
Optilab YDFA-r-40-S, Output power of 20W
Existing laser system
Simple Nd amplifier laser system without pulse shaping according to RF-Gun reviewer’s comment

## Slide 38: Strecher for Yb:YAG & Nd:YAG
2.0nm
 | Yb:YAG | Nd:YAG
Center wavelength | 1030 nm | 1064 nm
Gain spectrum width | ~2 nm | ~0.5nm
Distance of the Stretcher to 30 ps | 1.5 m | 6 m (×4)
Grating pair
30 ps
Gain spectrum of Yb:YAG
Distance of grating pair

## Slide 39: Nd:YAG DPSS module regenerative amplifier (experimental)
Test at A-1 underground using existing Yb-Fiber oscillator.
Regenerative amplifier
Oscillator

## Slide 40: Nd:YAG DPSS module / Northrup Grumann
円安 → ×3台 → 1200万円？

## Slide 41: Nd:YLF
蛍光寿命が長い ～485μs (Nd:YAG の倍)
発振波長が Yb に近い

## Slide 42: Improvement of fiber collimation
Fiber coupler
Cooling system
Fiber coupler in ERL
Drum Lens
PCF fiber damage is one issue => Improve the fiber collimation
Higher output pulse energy from fiber amplifier
=> Reduce number of stages of Yb:YAG multi-pass amplifier

## Slide 43: Thermionic gun
Thermionic Gun

## Slide 44
KL-A1-A
KL-A1-B
RF-gun
第2加速管
移相器・減衰器
（Pre-buncher）
現状
Electron gun rearrangement

## Slide 45
第1RF-gun
(第2RF-gun)
KL-A1-A
KL-A1-B
加速管2本
加速管１本
Buncher
Pre-buncher
WG 1本 新規製作
移相器・減衰器
（移設）
他のWG新規製作
又は流用
移相器・減衰器
（移設）
3dB DC
復元
WG 3本を
600mmつめる
(第2RF-gun用パルスベンドと予備スペース)
改造案2
S.Ohsawa
Electron gun rearrangement

## Slide 46: Preparation of Thermionic Gun
Under refurbishment
  Raise by 75cm not to conflict with straight RF-gun
    As well as angled RF-gun
  ~ Jun.2015.
May serve primary electron for positron generation
Thermionic gun
RF gun
Thermionic Gun

## Slide 47: Recent Works in Progress
Recent works
Thermionic gun
RF gun
#A1 region
#15 region
Iron shield

## Slide 51: Positron generator
Positron Enhancement

## Slide 52: Positron Generation
4-times more positron is required at SuperKEKB than KEKB
  Safety measure was taken after cable fire during the test of Flux Concentrator (FC)
  New components in 100-m capture section were tested in steps
  High voltage tests in tunnel in April
  Beam tests with electron in May
Positron Enhancement

## Slide 53: Positron generation for SuperKEKB
New positron capture section after target with Flux concentrator (FC) and large-aperture S-band structure (LAS)
Satellite bunch (beam loss) elimination with velocity bunching
Pinhole (2mm) for electrons beside target (3.5mm)
Beam spoiler for target protection
Flux
Concentrator
e+
5 nC injection e-
10 nC
primary e-
bridge coils
target
beam
hole
pulsed ST
DC QM
pulsed QM
side view
spoiler
solenoid
LAS Accel.
structure
positron
production
Target
Flux
Concentrator
Bridge
Coils
primary
e- beam
e+ beam
Positron Enhancement

## Slide 54: 陽電子捕獲部@Linacトンネル
この中にFC、標的
ブリッジコイルが
収められている。
電子ビーム
2014年4月に陽電子捕獲部
をビームラインの設置し、
5月より陽電子のコミッショニング
を開始した。
DCソレノイド
T.Kamitani
Positron Enhancement

## Slide 55: Positron from New Positron Capture Section
Generated positron ~0.1nC was transferred to the entrance of damping ring
With higher magnetic and electric field, 4-nC positron will be generated
Target shield (40cm x 6m long) will be finalized
Alignment will be improved 3mm  0.1mm
Horizontal
400 m
Vertical
Beam charge
Primary Electron
Beam position
Positron
W target
~0.1 nC
Flux concentrator and
following solenoid/quad
Large aperture S-band structure before solenoid & quad installation
Positron Enhancement

## Slide 56: (1) FC current
6kAより大きな電流値でもまだ saturationせずに順調に伸びていくと 期待される
新型電源では12kAまで上がる予定。
伝送路の inductance も下げる予定。
FC currentがzeroでも
ある程度のyieldがある
CT測定値
FCでの実電流値は1.07倍
plot上の6kA => 6.42 kA
T.Kamitani
Positron Enhancement

## Slide 57: 陽電子生成装置遮蔽体
遮蔽体の全体図
1-5 陽電子ターゲット下流の加速器上空に鉄を主体としたシールドを設置する．
S. Matsumoto
Positron Enhancement

## Slide 58
鳥瞰図：南東方向から見た図． トンネル1-5 陽電子ターゲットの60cm上流から６ｍの範囲に設置．3 月時の遮蔽鉄の厚みは200mm（図中の赤＋灰色部）．
通 路
雛 段
200mm角柱 6本が現場通路上ビームラインよりに設置される．鉛カーテンを設置すると、通路の幅が現場では1400mm．申請電流値に合わせて鉄の厚さを追加する予定．
Positron Enhancement

## Slide 59: Radiation Shield Construction
Radiation Shield
Flux Concentrator
#15 region
Iron shield

## Slide 60: Positron Generation
Installation of positron generator for SuperKEKB in April 2014 (Beamline construction since summer 2013) (positron target, spoiler, Flux Concentrator, bridge coils, LAS structures [x6], DC solenoids [16+13], e+/e- separator, quads [>90])
Commissioning of positron beam, observation of the first positron after reconstruction for SuperKEKB, further improvements expected
Oct.~Dec.2014 : Linac commissioning Jan.~Mar.2015 : Construction Apr.2015~ : Linac commissioning Feb.2016 : LER injection
T.Kamitani
 | Primary e- [nC] | Positron [nC] | Efficiency | Parameters
June 2014 | 0.6 | 0.12 | 20% | FC 6.4kA, Solenoids 370A, LAS capture field 10 MV/m
Specification (at SY2) | 10.0 | 5.0 | 50% | FC 12kA, Solenoids 650A, LAS capture field 14 MV/m
DR injection (2017?) |  | 4.0 | 40% | Energy spread acceptance 0.5%
x17
x42
x2.5
Positron Enhancement

## Slide 61: Schedule
Schedule

## Slide 62: Linac Schedule Overview
RF-Gun e- beam
commissioning
at A,B-sector
e- commiss.
at A,B,J,C,1
e+ commiss.
at 1,2 sector (FC, DCS, Qe- 50%)
e- commiss.
at 1,2,3,4,5 sector
non damped e+ commiss.
at 1,2, 3,4,5 sectors
e- commiss. at A→5 sectors
damped e+ commiss.
at 1→5 Qe+ = 1~4nC
e- commiss.
at A→5 Qe- = 1~5nC
: Electron
: Positron
: Low current electron
2 nC
Phase2
1 nC
Phase1
Location
↓
Time →
Phase1: high emittance beam for vacuum scrub
Phase2,3: low emittance beam for collision
Schedule
Without Top-up
4 nC
Phase3
Improved RF gun
Direct PF-AR BT
DR Commiss.
Staged Licenses
Low Emittance Beams
4+1 Ring
Injections

## Slide 63: Injector linac schedule
Feb.2016 – Jun.2016: Phase-1 commissioning
  Normal-emittance, 1nC/bunch electron/positron beams, without damping ring (DR)
    With combination of RF-gun and thermionic gun
    ex. Electron with 1nC RF-gun, Positron with ~6nC thermionic gun
      (depends on downstream configuration after DR delay affecting PF/PF-AR injections)
Jan.2017 – May.2017, Damping ring commissioning
  1nC – 2nC/bunch positron beam, to/from DR
Jun/Oct.2017 – Feb.2018, Phase-2 commissioning
  Low-emittance (20mm.mrad, 0.1%), 2nC electron/positron beams, with DR
    Low-emittance electron beam with RF-gun, 2nC
    Primary beam for positron with RF-gun or thermionic gun, 5nC
Oct.2018 – …, Phase-3 commissioning
  Low-emittance (20mm.mrad, 0.1%), High charge electron/positron beams, with DR
    Low-emittance electron beam with RF-gun, 4nC
    Primary beam for positron with RF-gun or thermionic gun, 10nC
Schedule

## Slide 64: Radiation protection licenses
Staged upgrade of beam limits
Final goal is 1250/625 nA before/after target
  Same as KEKB (with limited shields)
Applications
  Fall.2013. 10 nA at #28 dump, 1250 nA at #A2 dump
  Spring.2014. New utility rooms, 50 nA at #61 straight dump
  Feb.2015. 200 nA at #15 target
  Late 2015.(?) 800 nA at #15 target, 625 nA at #61
  Sometime 2016.(?) 1250 nA at #15 target
Schedule

## Slide 65: Summary
Steady progress towards first MR injection in 2015
Finished earthquake disaster recovery in 2014
Will make staged improvements up to Phase-III
Alignment: almost confident on the required precision (0.1-mm local, 0.3-mm global), need to maintain for longer term
RF gun: following recommendations at review meetings with commercial devices and Nd-based lasers
Thermionic gun: waiting to be commissioned
Positron generator: waiting for license test
Will balance between final beam quality and staged operation
Will select optimized route depending on available resources
Summary

## Slide 66
SuperKEKB
dual rings
Mt. Tsukuba
Injector Linac
PF-AR
PF
Thank you

## Slide 68: Design of a quasi traveling wave side couple RF gun
Normal side couple structure
Quasi traveling wave sidecouple structure
Quasi traveling wave side couple has stronger focusing field
RF gun for low-emittance electron
M. Yoshida et al

## Slide 69: Beam tracking simulation result
Emittance 5.5 mm-mrad
Size 0.4 mm
Energy spread 0.6%
Bunch shape
Gun Exit
5 nC
10 nC
RF gun for low-emittance electron
M. Yoshida

## Slide 70: RF-Gun comparison
DAW-type RF gun (90 MV/m, 5 mm-mrad, 3.2 MeV)
BNL-type RF gun (120 MV/m, 11.0 mm-mrad, 5.5 MeV)
Quasi traveling wave side couple RF gun (100 MV/m, 6mm-mrad, 13.5 MeV)
BNL (3.8 mm)
QTW (0.2 mm)
DAW (1.7 mm)
Beam Size
Quasi traveling wave side couple cavity
RF gun for low-emittance electron
M. Yoshida

## Slide 71: Ir5Ce Cathode
Quantum efficiency improvement
by Laser cleaning
×２０
LaB6
Carbon
Oxigen
SEM-EDX
Ir5Ce
Carbon
SEM-EDX
No oxidization
is observed
Lower energy density is enough to activate Ir5Ce
QE lifetime
QE Enhancement of IrCe cathode
Ir5Ce Cathode
RF gun for low-emittance electron
M. Yoshida

## Slide 72: Energy spread reduction using temporal manipulation
Gaussian
Square
5nC
10nC
15nC
20nC
20nC
15nC
10nC
5nC
5nC electron
15nC
Primary beam for positron production
t
t
Energy spread of 0.1% is required for SuperKEKB synchrotron injection.
RF gun for low-emittance electron
M. Yoshida

## Slide 73: Properties of laser medium
LD Pump
(808nm)
Nd:YVO4
Nd:YAG
SHG(532nm) 40%
FHG(266nm) 20%
5HG(213nm) 3%
τ～200μs, 40%
Nd-doped
Yb-doped
Ti-doped
LD Pump
(941/976nm)
Yb-glass
Yb:YAG
Yb:BOYS
1064nm
808nm
941/976nm
1040nm
τ～900μs, 40%
Pump
(808nm)
Nd:YAG
SHG
40%
1064nm
808nm
Ti:Sapphire
532nm
800nm
τ～3μs, 40%
τ=200μs, 40%
○ 4-state laser is easy to operate.
○ High power pump LD is available.
○ Large crystal is available
× Pulse width is determined by SESAM. (Gaussian)
○ Wide bandwidth => pulse shaping
○ Long fluorescent time => High power
○ Fiber laser oscillator => Stable
○ Small state difference
× ASE
× Absorption
○ Very wide bandwidth
○ High breakdown threshold
× Low cross section
× Short fluorescent time => Q-switched laser is required for pumping
Pump
TW laser is based on Ti-Sapphire
SHG(520nm) 40%
FHG(260nm) 20%
5HG(208nm) 3%
SHG(400nm) 40%
THG(266nm) 20%
FHG(200nm) 10%
Best for Linac RF-Gun
Nd laser system for 3-2 RF-Gun
Ti:Sapphire laser system.
RF gun for low-emittance electron

## Slide 74: A-1 RF gun results
A1 sector at KEK linac
x | y
32.7 ±3.1 mm-mrad | 10.7 ±1.4 mm-mrad
Q-scan emittance measurement
beam size measurement for Q-scan
RF gun
5.6 nC bunch charge was observed.

## Slide 75: Photo cathode RF gun development
5.6 nC / bunch was confirmed
Next step: 50-Hz beam generation & Radiation control
Quasi traveling wave side couple cavity
Cascaded frequency doublers
Yb fiber and Yb:YAG think disk laser
Ir5Ce photo cathode
RF gun for low-emittance electron
Part of multi-pass Amplifier

## Slide 76: GU_A1 Laser Configuration as of Nov.2014
RF gun for low-emittance electron

## Slide 77: GU_A1 Laser Configuration as of Dec.2014
RF gun for low-emittance electron

## Slide 78: Photo cathode RF gun improvement
Crucial for high-current low-emmittance beam
New Ir5Ce cathode and new cavity QTWSC were successful
Basic features were confirmed at 2 ~ 5 Hz
Expect beam parameter and stability performance at 50 Hz, with multi-pass amplifiers and cooling system
Resolved the issue of oscillator synchronization
Staged laser system improvements with beam measurement system
RF gun for low-emittance electron

## Slide 79: Pulse-to-pulse modulation
ARC
e– (2.5GeV, 0.2nC)
e− Gun
ARC
e+ Target
e+ (4GeV, 4nC)
e− Gun
PF Injection
SuperKEKB-LER Injection
e– (3.5GeV, 10nC)
ARC
e– (7GeV, 5nC)
e− Gun
SuperKEKB-HER Injection
ARC
e– (6.5GeV, 5nC)
e− Gun
PF-AR Injection
Damping ring
Event-based Control System
Every 20 ms
F.B
F.B
F.B
F.B
F.B
F.B
F.B
F.B
Four PPM virtual accelerators for SuperKEKB project maybe with additional PPM VAs for stealth beam measurements based on Dual-tier controls with EPICS and event-system